## NOTES
SPREADSHEET DOCUMENTATION
Author: | Tony Wahl
Publisher: | Bureau of Reclamation
Date of publication: | 2024-09-30
RISE Catalog record link: | https://data.usbr.gov/catalog/8049
item link (report): | https://data.usbr.gov/catalog/8049/item/128747
item link (data): | https://data.usbr.gov/catalog/8049/item/128751
Purpose: | Calibration data for Pitot tube used to measure boundary layer velocity profiles in hydraulic jacking research
Data source: | Hydraulics laboratory measurements
Period of record: | April 2024
Data description: | See final report ST-2024-22039 and the cited journal articles for more information.
Iterative calculation is required for this spreadsheet due to circular references. Excel does not always remember the iterative calculation setting. To re-enable iterative calculation, go to File / Options / Formulas and check the Enable iterative calculation box. (See screen shot at right.)

## Sheet1
Calibration Check of 0.052" Pitot tube on 2/23/2023
Nozzle diameter = |  | 8 | inches
| Area = | 0.34907 | ft2
Tank rim elev. on manometer board = |  |  |  | 1.57 | ft
Pitot diameter |  |  |  | 0.052 | inches
Point gage reading with dynamic port touching bottom edge of nozzle | [7 empty] | 14.19 | cm
Distance from centerline to center of Pitot dynamic port @ bottom position | [7 empty] | -3.974 | inches
Point gage at centerline | [7 empty] | 24.284 | cm
Max elevation | [7 empty] | 33.29 | cm
Interference zone (no measurements) from 17.21 to 21.40 cm
[12 empty] | Based on pitot diameter | No capillary rise correction, since both manometers are affected equally |  | Corrected for capillary rise in dynamic pressure manometer tube |  | Average of both methods
Venturi meter | Kventuri | Mercury column | Q | Vavg | V2/(2g) | Point gage elevation | Distance from boundary | Static pressure (Pitot) | Dynamic pressure (Pitot) | Tank water level below rim, inches | Static pressure, calculated from tank water level | Re | DPPitot-static | C | DPPitot-Tank | C | Cavg | SQRT(Cavg) | r | v | r* | v*
|  | mm Hg | cfs | ft/s | ft | cm | % | ft | ft | inches | ft | - | ft | - | ft | - | -
NE 8" | 0.18565 | 417.29 | 3.7924 | 10.8644 | 1.8346 | 24.28 | 0.99961 | 1.328 | 3.315 | -3.25 | 1.2992 | 4441.4299 | 1.987 | 1.0831 | 1.9954 | 1.0877 | 1.0854 | 1.0418 | 0.3332 | 11.3188 | 0.99961 | 1.0418
NE 8" | 0.18565 | 417.29 | 3.7924 | 10.8644 | 1.8346 | 24 | 0.97205 | 1.328 | 3.315 | -3.25 | 1.2992 | 4441.4299 | 1.987 | 1.0831 | 1.9954 | 1.0877 | 1.0854 | 1.0418 | 0.32402 | 11.3188 | 0.97205 | 1.0418
NE 8" | 0.18565 | 417.29 | 3.7924 | 10.8644 | 1.8346 | 21.5 | 0.72599 | 1.333 | 3.292 | -3.25 | 1.2992 | 4441.4299 | 1.959 | 1.0678 | 1.9724 | 1.0752 | 1.0715 | 1.0351 | 0.242 | 11.2461 | 0.72599 | 1.0351
NE 8" | 0.18565 | 417.29 | 3.7924 | 10.8644 | 1.8346 | 17.2 | 0.30276 | 1.355 | 3.212 | -3.25 | 1.2992 | 4441.4299 | 1.857 | 1.0122 | 1.8924 | #N/A | 1.0122 | 1.0061 | 0.10092 | 10.9307 | 0.30276 | 1.0061
NE 8" | 0.18565 | 417.29 | 3.7924 | 10.8644 | 1.8346 | 14.5 | 0.037012 | 1.35 | 2.855 | -3.25 | 1.2992 | 4441.4299 | 1.505 | 0.82036 | 1.5354 | 0.83695 | 0.82865 | 0.9103 | 0.012337 | 9.8899 | 0.037012 | 0.9103
NE 8" | 0.18565 | 417.29 | 3.7924 | 10.8644 | 1.8346 | 25 | 0.92952 | 1.322 | 3.305 | -3.25 | 1.2992 | 4441.4299 | 1.983 | 1.0809 | 1.9854 | 1.0822 | 1.0816 | 1.04 | 0.30984 | 11.2989 | 0.92952 | 1.04
NE 8" | 0.18565 | 417.29 | 3.7924 | 10.8644 | 1.8346 | 26 | 0.8311 | 1.32 | 3.305 | -3.25 | 1.2992 | 4441.4299 | 1.985 | 1.082 | 1.9854 | 1.0822 | 1.0821 | 1.0403 | 0.27703 | 11.3017 | 0.8311 | 1.0403
NE 8" | 0.18565 | 417.29 | 3.7924 | 10.8644 | 1.8346 | 27 | 0.73267 | 1.32 | 3.295 | -3.25 | 1.2992 | 4441.4299 | 1.975 | 1.0766 | 1.9754 | 1.0768 | 1.0767 | 1.0376 | 0.24422 | 11.2732 | 0.73267 | 1.0376
NE 8" | 0.18565 | 417.29 | 3.7924 | 10.8644 | 1.8346 | 28 | 0.63425 | 1.32 | 3.282 | -3.25 | 1.2992 | 4441.4299 | 1.962 | 1.0695 | 1.9624 | 1.0697 | 1.0696 | 1.0342 | 0.21142 | 11.2361 | 0.63425 | 1.0342
NE 8" | 0.18565 | 417.29 | 3.7924 | 10.8644 | 1.8346 | 29 | 0.53582 | 1.316 | 3.252 | -3.25 | 1.2992 | 4441.4299 | 1.936 | 1.0553 | 1.9324 | 1.0533 | 1.0543 | 1.0268 | 0.17861 | 11.1556 | 0.53582 | 1.0268
NE 8" | 0.18565 | 417.29 | 3.7924 | 10.8644 | 1.8346 | 30 | 0.4374 | 1.315 | 3.225 | -3.25 | 1.2992 | 4441.4299 | 1.91 | 1.0411 | 1.9054 | 1.0386 | 1.0399 | 1.0197 | 0.1458 | 11.0789 | 0.4374 | 1.0197
NE 8" | 0.18565 | 417.29 | 3.7924 | 10.8644 | 1.8346 | 31 | 0.33897 | 1.305 | 3.178 | -3.25 | 1.2992 | 4441.4299 | 1.873 | 1.021 | 1.8584 | 1.013 | 1.017 | 1.0085 | 0.11299 | 10.9563 | 0.33897 | 1.0085
NE 8" | 0.18565 | 417.29 | 3.7924 | 10.8644 | 1.8346 | 32 | 0.24055 | 1.3 | 3.135 | -3.25 | 1.2992 | 4441.4299 | 1.835 | 1.0002 | 1.8154 | 0.98957 | 0.99491 | 0.99745 | 0.080182 | 10.8367 | 0.24055 | 0.99745
NE 8" | 0.18565 | 417.29 | 3.7924 | 10.8644 | 1.8346 | 33.29 | 0.11358 | 1.275 | 2.99 | -3.25 | 1.2992 | 4441.4299 | 1.715 | 0.93483 | 1.6704 | 0.91054 | 0.92268 | 0.96056 | 0.03786 | 10.436 | 0.11358 | 0.96056
[11 empty] | FLIPPED AXES… (and dimensional, v as function of r) | [9 empty] | Fitted Line from boundary to center
[21 empty] | X | Y | x* | y*
CURVE FIT PARAMETERS FOR VELOCITY PROFILE | [20 empty] | 0 | 0 | 0 | 0
| a = | 11.9215 | [18 empty] | 1e-18 | 2.2194 | 3e-18 | 0.20428
| b = | 0.040561 | [18 empty] | 1e-15 | 2.937 | 3e-15 | 0.27034
[21 empty] | 1e-12 | 3.8868 | 3e-12 | 0.35776
Integrated Discharge function for the whole pipe | [20 empty] | 1e-09 | 5.1437 | 3e-09 | 0.47345
| R | 0.33333 | ft | [17 empty] | 1e-06 | 6.8071 | 3e-06 | 0.62655
| Q | 3.7489 | ft3/s | [17 empty] | 5e-06 | 7.2663 | 1.5e-05 | 0.66882
[21 empty] | 2.5e-05 | 7.7565 | 7.5e-05 | 0.71394
|  |  | Error from venturi meter = | -1.148% | [16 empty] | 0.000125 | 8.2797 | 0.000375 | 0.7621
[21 empty] | 0.000625 | 8.8383 | 0.001875 | 0.81351
[21 empty] | 0.003125 | 9.4345 | 0.009375 | 0.86839
[21 empty] | 0.00625 | 9.7035 | 0.01875 | 0.89315
[21 empty] | 0.0125 | 9.9802 | 0.0375 | 0.91861
[21 empty] | 0.025 | 10.2648 | 0.075 | 0.94481
[21 empty] | 0.05 | 10.5575 | 0.15 | 0.97175
[21 empty] | 0.1 | 10.8585 | 0.3 | 0.99946
CURVE FIT PARAMETERS FOR DIMENSIONLESS VELOCITY PROFILE | [9 empty] | Fitted Line from boundary to center | [10 empty] | 0.2 | 11.1681 | 0.6 | 1.028
| a = | 1.1014 | [7 empty] | X | Y | [9 empty] | 0.33333 | 11.402 | 1 | 1.0495
| b = | 0.081123 | [7 empty] | 0 | 0 | [11 empty] | 1.0495
[10 empty] | 1e-18 | 0.038171 | [11 empty] | Velocities are about 5% high at centerline
[10 empty] | 1e-15 | 0.06685
[10 empty] | 1e-12 | 0.11708
[10 empty] | 1e-09 | 0.20504
[10 empty] | 1e-06 | 0.35909 | [12 empty] | ATTEMPT TO FIT CURVE BETTER WITH A CUSTOMIZED FUNCTION…
[10 empty] | 0.0001 | 0.52174 | [12 empty] | c | 0.041892
[10 empty] | 0.001 | 0.62889 | [12 empty] | a | 1.0412 |  | S(err)2 | 2.2022e-05
[10 empty] | 0.005 | 0.7166 | [12 empty] | b | -0.040302
[10 empty] | 0.025 | 0.81655 | [13 empty] | y = a(x)^(c+bx)
[10 empty] | 0.125 | 0.93043 | [13 empty] | Fitted line |  | Data…................
[10 empty] | 0.19 | 0.96258 | [13 empty] | r* | v* | r* | v* | v*predicted
[10 empty] | 0.2888 | 0.99584 | [13 empty] | 0 | 0 | 0.99961 | 1.0418 | 1.0412
[10 empty] | 0.43898 | 1.0302 | [13 empty] | 1e-10 | 0.39685 | 0.97205 | 1.0418 | 1.0412
[10 empty] | 0.66724 | 1.0658 | [13 empty] | 1e-09 | 0.43704 | 0.72599 | 1.0351 | 1.037
[10 empty] | 1 | 1.1014 | 1.1014 | [12 empty] | 1e-08 | 0.4813 | 0.30276 | 1.0061 | 1.0049
[12 empty] | C values are about 10% high at centerline because they relate to V2 | [12 empty] | 1e-07 | 0.53004 | 0.037012 | 0.9103 | 0.9114
[25 empty] | 1e-06 | 0.58371 | 0.92952 | 1.04 | 1.0409
[25 empty] | 1e-05 | 0.64282 | 0.8311 | 1.0403 | 1.0396
[25 empty] | 7.5e-05 | 0.69946 | 0.73267 | 1.0376 | 1.0372
[25 empty] | 0.000375 | 0.74831 | 0.63425 | 1.0342 | 1.0335
POINTS THAT DON'T FIT VERY WELL (Outliers regardless of whether static port or physical tank-level measurement is used) | [24 empty] | 0.00125 | 0.78719 | 0.53582 | 1.0268 | 1.0281
NE 8" | 0.18565 | 417.29 | 3.7924 | 10.8644 | 1.8346 | 16 | 3.2614 | 1.342 | 3.205 | -3.25 | 1.2992 | 4441.4299 | 1.863 | 1.0155 | 1.8854 | 1.0277 | [8 empty] | 0.00625 | 0.84289 | 0.4374 | 1.0197 | 1.0206
NE 8" | 0.18565 | 417.29 | 3.7924 | 10.8644 | 1.8346 | 15 | 3.6551 | 1.325 | 3.125 | -3.25 | 1.2992 | 4441.4299 | 1.8 | 0.98116 | 1.8054 | 0.98412 | [8 empty] | 0.0125 | 0.86853 | 0.33897 | 1.0085 | 1.0099
NE 8" | 0.18565 | 417.29 | 3.7924 | 10.8644 | 1.8346 | 19 | 2.0803 | 1.335 | 3.138 | -3.25 | 1.2992 | 4441.4299 | 1.803 | 0.9828 | 1.8184 | 0.99121 | [8 empty] | 0.025 | 0.89547 | 0.24055 | 0.99745 | 0.99455
NE 8" | 0.18565 | 417.29 | 3.7924 | 10.8644 | 1.8346 | 33 | 3.4315 | 1.298 | 3.082 | -3.25 | 1.2992 | 4441.4299 | 1.784 | 0.97244 | 1.7624 | 0.96068 | [8 empty] | 0.05 | 0.92399 | 0.11358 | 0.96056 | 0.96006
[25 empty] | 0.1 | 0.95431
[25 empty] | 0.15 | 0.97278
[25 empty] | 0.2 | 0.98606
[25 empty] | 0.25 | 0.99631
[25 empty] | 0.3 | 1.0045
[25 empty] | 0.35 | 1.0113
[25 empty] | 0.4 | 1.0169
[25 empty] | 0.5 | 1.0257
[25 empty] | 0.6 | 1.0319
[25 empty] | 0.7 | 1.0362
[25 empty] | 0.8 | 1.039
[25 empty] | 0.9 | 1.0406
[25 empty] | 1 | 1.0412
The Pitot tube was used to traverse the height of the 8-inch ASME flow nozzle. Mechanical interference in the mounting system prevented measurements in a band below centerline, but measurements were made in the symmetric zone above centerline. No measurements could be made near the top edge due to the static port striking the edge of the nozzle. The legs of the Pitot tube projected about 1/4-inch upstream from the downstream edge of the nozzle. Measurements of dynamic pressure and static pressure were made on a nearby manometer board. In case the static port pressures were inaccurate, measurements of the tank water level were also made by tape measure. The water level in the tank is quite turbulent for larger flow rates, so the tank-level measurements are only approximate.
This gives a great fit to the observed data but is not something that can be directly integrated to get the total discharge.
Note on Q22: Wahl, Tony L: This point is an outlier, but the measurement from the static port seems good.